	SCR number and title
	Project

	0003496: Head: Base: Parameter Name Differed Between PharmGuard Override and Normal Therapy Change
	Medfusion 4000


	Test Pump Serial #
	Version/Build Number of Code
	configuration name
	configuration guid
	configuration date/time

	Tested on Real pump serial no DULUTH01.
	26555
	ParameterNameAnomaly.medcfg
	N/A
	N/A


	Step
	Input
	Expected Output
	Actual Output (“as expected” or description)
	Pass/Fail

	1.
	Initial conditions:
Load the ParameterNameAnomaly.medcfg
configuration.
Load a BD 60 ml syringe.

Press “POWER” button on the pump to start.
	“PRESS THE NUMBER TO SELECT THE PROFILE” screen is displayed after the Startup.
	As expected
	Pass

	2.
	Select: “CARDIAC".

Select “Miscellaneous”.

Select “Adenosine (Adenocard) 3mg/ml”. 
	“CONFIRM LIB SELECTION?” screen is displayed on the pump.
	As expected
	Pass

	3.
	Press “YES”

Select BD and Press Enter
	“ENTER PATIENT WEIGHT” screen is displayed on the pump
	As expected
	Pass

	4.
	Enter 10 as a Patient weight and Press enter.
	“ENTER DOSE RATE” screen is displayed on the pump.
	As expected
	Pass

	5.
	Enter 150 as Dose Rate, silence alarm, and press override.
	“ENTER VOLUME LIMIT RATE” screen is displayed on the pump.
	As expected
	Pass

	6.
	Enter a volume limit (20ml)

Confirm Settings.

Start Infusion.
	Run screen is displayed on the pump
	As expected
	Pass

	7.
	Press "CHG Dose".

Enter 140 for dose.

Press Enter.
	Run screen is displayed on the pump
	As expected
	Pass

	8.
	Stop the infusion.

Enable the external communications on pump in Biomed and learn the history with Toolbox.

Observe the Parameter in PharmGuard override event on the Toolbox
	 “Dose” is displayed as a parameter on the Toolbox for the Pharmguard override event.
	As expected
	Pass

	9.
	Observe the parameter in the Biomed mode history display(Biomed->Utilities->Viewinfusionhistory) for pharmguard data override event.
	“Dose” is displayed as a parameter on the pump for the Pharmguard override event.
	As expected
	Pass


Comments:






































  Performed by: 
Ashok .K
 Date:  
16th Dec 09
   



